# APPENDIX A: FORCE PLATE Pilot Studies

## Transmission of Shear Forces Along Artificial Turf

The original strategy for this project was to combine all the outputs of the three available force plates in the laboratory (Figure 4.1) to calculate the external ground loadings applied to the subject.  It would have enabled a larger area for the subjects to perform movements.  An experiment was conducted to assess the ability of the turf to transmit forces horizontally along the turf.  This was done to decide how to the lay the turf over the force plates, in terms of practicality for calibration of the lab and force plate accuracy. Before this experiment, there were three options or configurations:

- Lay the turf over the three force plates in one continuous strip

- Place a smaller, rectangular strip of turf over the three force plates

- Isolate each force plate by cutting out 60x40cm of turf.

However, any transmission of shear forces along the turf may result in miscalculation of subject segmental joint moments during the inverse dynamics process. If there is significant transmission of horizontal forces, loads may be measured on a force plate when there is no direct loading on the force plate by the subject.

The aim of this experiment was to assess which configuration of turf layout resulted in the minimal transmission of horizontal force.

The previously described turf layout configurations were assessed.  For configuration 1, an 8m roll of artificial turf was laid out over a 1x1.5m section of rubber underlay, which was placed over the three force plates. For configuration 2, a 1.7x2m section of artificial turf was placed over the same section of rubber underlay on top of the force plates.  For configuration 3, the isolation of the force plates was mimicked by not placing any turf or underlay over the force plates. This configuration also acted as a control for this experiment.  For configurations 1 and 2 the positions of the force plates directly underneath were taped on the top surface of the turf.

Figure 4.1 Force plate layout in biomechanics lab

One subject performed 4 types of movement of each turf configuration. The four movements were straight running, running and stopping on the force plate, a running jump, and a running 45º turn to the right by pushing off from the subject left foot.  All movements were performed at maximal speed for the given movement.  The subject was instructed to target the force plate (FP B) in order to perform the respective movement, and not touch the other two force plates (FP A and FP C).

Forces in the each direction (Fx, Fy, Fz) were measured on all three force plates. The amount of force transfer in one direction was measured by calculating the average of the absolute peak readings from FP A and FP C as a percentage of the of the peak force on FP B.  Due to the origin of FP B being located midway between the origins of FP A and C in the z-direction, it was assumed that any force transference would be approximately equal to the other two force plates.

Any zero offsets that were present were removed through the Vicon software.

## Results

| Configuration | Running | Stopping | Jumping | 45º turn | Mean |
| --- | --- | --- | --- | --- | --- |
| Roll | 1.64 | 1.60 | 3.19 | 2.09 | 2.13 |
| Section | 1.95 | 1.35 | 2.82 | 2.32 | 2.11 |
| Control | 0.23 | 0.60 | 0.33 | 0.38 | 0.39 |

Table 4.1Fx (A-P) transfer as a percentage of the peak force (Fx) on FP B

| Configuration | Running | Stopping | Jumping | 45º turn | Mean |
| --- | --- | --- | --- | --- | --- |
| Roll | 1.40 | 3.99 | 2.88 | 2.00 | 2.57 |
| Section | 1.65 | 1.54 | 1.88 | 2.32 | 1.85 |
| Control | 0.19 | 2.13 | 0.66 | 0.36 | 0.84 |

Table 4.2 Fz (M-L) transfer as a percentage of the peak force (Fz) on FP B

| Configuration | Running | Stopping | Jumping | 45º turn | Mean |
| --- | --- | --- | --- | --- | --- |
| Roll | 0.13 | 0.17 | 0.11 | 0.15 | 0.14 |
| Section | 0.14 | 0.24 | 0.19 | 0.23 | 0.20 |
| Control | 0.13 | 0.20 | 0.11 | 0.15 | 0.15 |

Table 4.3 Fy (vertical) transfer as a percentage of the peak force (Fy) on FP B

Trials were discarded if it was obvious from the data that the subject had touched FP A and FP C during the movement.

It is clear from Table 4.1andTable 4.2 that the presence of a continuous section or roll of turf placed over the three force plates allows the transfer of shear force. Although this transfer may only be approximately 2% of the peak force, this would represent an additional error in the measurement process.

The control configuration had a force transfer of less than 0.85% of the peak force in each direction. This demonstrated the minimal cross talk that occurs between the isolated force plates for the movements performed.

The 2.13% cross talk on the Fz (mediolateral) during stopping in the control condition stands out as rather high, compared to the other results in this condition. Section 4.2 assessed this problem further by examining the cross talk between the force plates during similar high frequency movements.

## Force Plate Saturation

Only a few trials were conducted in this experiment as it became apparent that the channels on FP B were becoming saturated during some of the movements involving high impact.

Separate trials involving high impact movements on FP A and FP C revealed that saturation occurred on these plates also at high impact (Table 4.4).  This meant the gain settings were too low and would have to be increased if ground reaction forces of higher magnitudes were to be measured.

With regards to the effect of this saturation on the results above, the percentage cross talk of peak load may have been slightly exaggerated where saturation of the peak force occurred.  However, after extrapolation of the graphs of those trials where saturation occurred, the peak forces were not reduced greatly.

| Platform | Fx (N) | Fy (N) | Fz (N) |
| --- | --- | --- | --- |
| A | 1284 | 5184 | Not reached |
| B | 548 | 1984 | 508 |
| C | 510 | 2074 | 517 |

Table 4.4 Maximum force at saturation level

Force Plate A:

The nominal force range for X and Y (horizontal) forces was required to be increased from ~500N to ~1000N per channel.  Therefore, the range on Force Plate A’s charge amplifier was set to 10,000pC:

1000N (nominal force) x 8pC/N (sensitivity) = 8000pC (RANGE)

Therefore the next highest measuring range must be selected, i.e. 10,000pC.

The range currently selected for the Z (vertical) force is 5000pC, which allows a force range of ~1300N per channel or ~5200N in total:

1300N (nominal force) x 3.8pC/N (sensitivity) = 4940pC (RANGE per chan.)

Force Plates B & C:

The transducer sensitivity range is an indicator guide to the amplitude of the range setting and it is selected according to the decade on the sensitivity value of the transducer. (eg if the sensitivity value is 7.9 pC/N, the transducer sensitivity range was set at 1…..11).

The nominal force range for the summated X and Y (horizontal) forces  (top 4 charge amplifiers) require to be increased from ~500N to ~1000N per channel. The units of gain for these charge amplifiers were in mech units/V (N/V): The gain was selected as follows:

1000N (nominal force) / 10 V (voltage range) = 100 N/V (GAIN)

The nominal force range for the summated Z (vertical) forces  (bottom 4 charge amplifiers) require to be increased from 2000N to 5000N per channel. The units of gain for these charge amplifiers are in mech units/V (N/V): The gain was selected as follows:

5000N (nominal force) / 10 V (voltage range) = 500 N/V

The calibration scale factor within the Vicon software (Analogue Setup, using a .car file) to calculate the force plate outputs into Newtons of force was adjusted accordingly for the new gain settings.

## Force Plate Cross Talk Analysis

### Introduction

From the results from the experiments examining the transmission of shear force, it became apparent that the ‘cross talk’ observed may have come from two sources or a combination of both:

- Transmission of shear forces along the turf.

- Transmission of forces caused by resonance through the force plate support structure.

Although during the previous experiments, the cross-talk in the control condition (no turf – isolated force plates) there was minimal cross-talk from FP B to FP A and C, there was one instance when the mean force transfer was 2.13% of a peak load of 203N in the mediolateral direction during a running-stop movement.

It was decided to investigate this further by conducting trials on all three force plates and measuring the amount of cross talk to the other two plates.  This experiment would allow a better understanding of the amount of cross-talk that occurs through the resonance of the force plate support system during high impact movements.  During subject testing, the 6 outputs of all three force plates were to be combined to produce a single set of ground loading outputs.  Significant cross talk on the force plates that may not be loaded during a trial may introduce significant error into these combined force and moment calculations.

### Method

One subject conducted ten trials on each of the three force plates.  The movement performed was similar to that of the take-off technique for a long jump.  The subject ran towards the respective force plate in a straight line and planted the take-off foot wholly on the centre of the force plate.  For each force plate, five trials were conducted in the lab x-direction and five trials were conducted in the lab y-direction. The subject varied the speed of run up and force that the take-off foot was placed on the force plate for each trial.

Trials were discarded if it was obvious that the take-off foot was not placed solely on the respective force plate tested.  This was checked by analysing the position of the reconstructed ground reaction vector on the force plate through the Vicon software.  If the ground reaction vector was located towards the centre of the plate, the trial was accepted.  If the ground reaction vector was located towards the edge of the force plate, the trial was discarded.

For each trial, all three force plates were sampled at 200Hz.  The gain settings were increased for FP B and FP C to 100 mu/V for the horizontal channels and to 500mu/V for the vertical channels to allow a greater range of forces to be recorded. The calibration scale factors within the Vicon software was appropriately amended.

The force plate output was measured through the Vicon system. This data was converted into an ASCII format, which was then inputted into MATLAB program for analysis.  This calculated the peak amount of cross talk on the two plates not loaded by the subject as a percentage of the peak force applied to the loaded force plate.

Cross talk on Fx channels during movement in z-direction

Cross talk on Fz channels during movement in x-direction

Results

Figure 4.2 Grouped cross-talk data from trials on FP A

Figure 4.2 shows the grouping of all peak loads in all directions (Fx, Fy and Fz) on FP A and their corresponding relative cross talk percentages measured on FP B and FP C.  FP1, FP2 and FP3 corresponds to Force Plates A, B and C, respectively.  For all trials conducted on FP A, the cross talk to the other two plates was less than 4% of the peak load.  The majority of the cross talk values are below 1% of the peak load.  The cross- talk values that were greater than 1% fall into two categories. The greatest cross talk values on the Fx channels on FP B and FP C occurred during movements in the z-direction.  The greatest cross talk values on the Fz channels on FP B and FP C occurred during movements in the x-direction.

Cross talk on Fx channels during movement in z-direction

Cross talk on Fz channels during movement in x-direction

Figure 4.3 Grouped cross talk data for trials on FP B

Figure 4.3 demonstrates that the majority of cross talk values are below 5% of the peak directional load. It is noticeable that cross talk greater than 5% was measured only on FP A.  These large cross talk percentages ranged from 5.2% to 16.8% of the peak directional load. Figure 2 also show that these values over 5% fall into two categories. The greatest cross talk values on the Fx channels on FP B and FP C occurred during movements in the z-direction.  The greatest cross talk values on the Fz channels on FP B and FP C occurred during movements in the x-direction.

Figure 4.4 Grouped cross talk data for trials on FP C

Similarly, most of the trials on FP C produced less than 5% cross talk to the other two force plates (Figure 4.4).  For the cross talk values greater than 5%, the same neat categories observed during trials on FP A and FP B were not apparent on trials on FP C.  A small group of cross talk values, ranging between 9–17% of the peak directional load (Fx: 206-458N), was observed on the Fx channels of both FP A and FP C during movements occurring in the x-direction.

However, during one single trial (CX3) when the subject landed from a jump on one foot on FP C, a cross talk ranging between 22% and 55% of their respective direction peak loads were measured.  The vertical force (4582N) produced during this trial peaked close to the maximum force before saturation of the charge amplifier would occur (5000N).  This near overloading of the charge amplifier may have been the reason for the rapid spike seen on the vertical force measurements on FP A and B (Error: Reference source not found).  Similar spikes were also observed as cross talk on the horizontal channels of the other two force plates. (Error: Reference source not found and Figure 4.7).

Figure 4.5 Vertical forces during landing on FPC (FP3)

Figure 4.6 Anteroposterior forces during landing on FP C (FP3)

Figure 4.7 Mediolateral forces during landing on FP C (FP3)

(a)

(b)

An attempt was made to filter out these unwanted cross talk values on the other two force plates.  A residual analysis (Winter, 1990), using a two-way 2nd order Butterworth filter, and analysis of the frequency spectrum of the vertical force from FP C (Figure 4.8a) and the vertical force from FP A (Figure 4.8b) suggested a cut-off frequency of 25-30Hz should be used.

Figure 4.8 Frequency spectrums of the a) vertical force output from FPC and b) vertical force output from FPA during landing on FPC

As a result of the problems identified with force transference along the artificial surface and cross talk between force plates during high impact loadings, it was decided to limit the use of the number of force plates to one.  However, this would reduce the measurement area in which the subjects could perform turns.  It was decided that Force Plate A would be used as this was the newer force plate and possessed the greatest natural frequency (800Hz). A single section of artificial turf and underlay (if required), the same area as the force plate (600x400mm), would be securely attached to the force plate. This would isolate this section of turf from the rest of the artificial turf.

## Force Plate Calibration Check

### Introduction

Calibration checks pf measuring equipment is necessary to ensure that the outputs of the device reflect the actual physical loads or parameters it is measuring. An in situ calibration of the force plates located in the Biomechanics Lab, Bioengineering Unit was conducted.  The force plate used on this study was a Kistler model 9281B11.  The technical specifications for this type of force plate are presented in Table 4.5.

| Model | Parameter | Specification (%FSO) |
| --- | --- | --- |
| FP A (9281B11) | Linearity | ±0.5 |
|  | Hysteresis | <0.5 |
|  | Cross-talk (Fz→Fx/y) | ±1 |
|  | Cross-talk (Fx↔Fy) | ±1 |
|  | Cross-talk (Fx/y→Fz) | ±2 |

Table 4.5 Technical specifications for Kistler force plates

This section describes the calibration check of the whole load analysis system. This included the force plate, associated charge and summing amplifiers, and Vicon motion analysis system.  The calibration check include the analysis of vertical and shear loading and the assessment of linearity, hysteresis and cross talk.

### Method

Before calibration of the force plates took place, it was reported that the force plate rocked slightly on their mounting blocks.  Replacing and reinforcing the shims under the mounting blocks that were used to stabilise the force plates corrected this problem.

The method used in this calibration check has previously been reported by Hall et al. (1996).  Three procedures were performed.  Firstly, a dead weight test was performed in which 10kg masses were placed on the centre of force plate. The masses were placed on top of each other until a total load of 1079.2N (110kg) was achieved.  Loads were than taken of one-by-one to test for hysteresis.

A point loader, which allowed the transference of a vertical load to a small steel ball, was used to check the calibration of vertical loading (Vicon system Z-axis) at nine recorded positions on the force plate surface. This method also allowed the applied moments around the force plate horizontal axes to be checked.

To check the calibration during shear loading, a rig was used which allowed the transference of a hanging vertical load, through a system of cables and pulleys to a latch plate.  The latch plate was located at the edge of the force plate and lay flush with the force plate surface.  This allowed a shear force to be applied to the force plate.  Shear force were applied along the force plate’s anteroposterior axis (Vicon system X-axis) and mediolateral axis (Vicon system Y axis) in positive and negative directions.  To be able to check the measurement of applied moments around the force plate’s vertical axes, the latch plate was offset from the centre of the force plate at recorded positions.

Due to safety and the limitations of the vertical and shear loading rigs, the force plates were not calibrated up to its full measuring range (0 – 10kN). Therefore, a calibrated part measuring range was used (0 – 490.5N). All results are presented as a percentage of the respective calibrated part measuring range full-scale output (%prFSO). All charge amplifiers were ‘zeroed’ prior to commencement of the calibration process.

The parameters calculated in this study are described as follows.

a) Mean Difference Sensitivities (MDS): Calculated from the mean error (minus the ‘zero load offset’) as a percentage of the applied load for N locations and load levels:

MDS=measured load−applied loadapplied load⋅100N		(Eq.  4.)

b) Linearity: expressed as ‘the deviation of a calibration curve from a specified straight line’ (ISA). Calculated as the mean  (minus the ‘zero load offset’) largest difference from the applied load during each calibration trial in each position, as a percentage of the part range FSO.

c) Hysteresis: expressed as ‘the biggest difference in the output signal for any value of the measured variable within the specified range when this value is reached with increasing or diminishing measurand’ (ISA).

d) Cross talk: expressed as ‘a signal at the output of a transducer caused by a measured variable acting on this transducer but not allocated to this particular output’ (Kistler).  This was calculated as the mean cross talk output (minus the ‘zero load offset’) from each position at an applied load of 490.5N (1082.5N for the dead weight trials). Cross talks for forces are calculated as a percentage of the applied load.

e) Centre of pressure calculation error: the mean error of the centre of pressure calculations for the x and y coordinates was assessed.  This was conducted on the force plate for the vertical point loadings for each of the nine positions at three loads: 0N, 10N and 50N.

### Calibration check results

| Applied Load | MDS (%) | Linearity (%prFSO) | Hysteresis (%prFSO) |
| --- | --- | --- | --- |
| Fz (dead weight) | -0.23 | 0.57 | 0.12 |
| Fx Fy Fz Mx My Mz | -0.49 -1.10 -1.04 -3.44 2.47 -4.60 | 0.95 1.16 0.89 -3.62 2.17 -4.71 | 0.44 0.46 0.34 0.19 0.46 0.30 |

Table 4.6 Calibration check results

| Applied Load | Parameter | Crosstalk |
| --- | --- | --- |
| Fz (dead weight) | Fz→Fx | 0.12 |
|  | Fz→Fy | 0.2 |
| Fz (point) | Fz→Fx | -0.440 |
|  | Fz→Fy | -1.150 |
| Fx | Fx→Fy | -0.307 |
|  | Fx→Fz | 0.507 |
| Fy | Fy→Fx | -0.398 |
|  | Fy→Fz | 0.211 |

| Coordinate |  | 0N | 10N | 50N |
| --- | --- | --- | --- | --- |
| X | mean error | 58.34 | 4.59 | 3.28 |
|  | (s.d.) | (114.69) | (48.24) | (10.66) |
| Y | mean error | 18.64 | 8.39 | 1.16 |
|  | (s.d) | (85.18) | (24.06) | (2.20) |

Table 4.7 Overall error means for COP calculation at three vertical loads

Figure 4.9 COP error (x coordinate)

Figure 4.10 COP error (y coordinate)

### Discussion

Hysteresis and crosstalk results from Table 4.6 show that they fall within Kistler’s specifications. Results for MDS and linearity for applied forces are slightly higher than Kistler’s specifications. However, they are either below or just slightly higher than ±1%, which is acceptable.

MDS and linearity results for the applied moments Mx, My and Mz are noticeably greater than Kistler’s specifications. The greatest was recorded on the Mz channel with a MDS of –4.6% and a linearity of –4.7% prFSO. These erroneous results may have been due to errors in the centre of pressure calculations at low load levels (Figure 4.9, Figure 4.10 and Table 4.7).  Furthermore, in the case of Mz, a free moment may have been introduced caused by the unobserved rotating of the weights on the top of the steel ball during the vertical loading trials.

## Assessment of the Inertial Properties of Artificial Surfaces

### Introduction

With the increase in the use of artificial turf in sports, there is a need to measure the human/surface interactions accurately in order to identify possible mechanisms that could lead to injury. Previous assessments of artificial turfs utilised mechanical rigs that apply loads to the surface that are not biomechanically relevant (Nigg, 1990). Subject testing of artificial turf usually employs a force plate to measure the loadings applied by the subject during dynamic movements. The turf is attached to the top surface of the force plate, which measures the loads transmitted from the subject through the artificial turf. Under high dynamic loading the artificial surface may exhibit an inertial effect, which could significantly affect the measurement of the applied load by the force plate:

Ffp = Fapp - mtat						(Eq. 4.)

Ffp   = force measured by the force plate

Fapp  = force applied

mtat = inertial force of the turf

There has not been any previous research that has examined this possibility. The aim of this experiment was to assess the accuracy of using a force plate to measure the dynamic loading actions at the human/surface interface on artificial turfs.

### Methods

ARTIFICIAL TURF

A test rig (Figure 4.11) was used to measure the vertical force produced from dynamic loads (Praagman, 1997). The rig was constructed to transmit vertical loads to a single-axis Kistler 9321B force transducer  (‘Force-link’) placed on top of the artificial turf, which was attached to a force plate (Kistler 9281B). The rig was designed to minimise any friction and prevent the transmission of any horizontal loads. The aluminium plates were hinged to a vertical steel support plate, which was bolted to the floor surrounding the force plate to provide stability.

Figure 4.11 Schematic of inertia testing rig

Three artificial turfs were tested: Turf 1 (Astroturf®: short pile, sand infill); Turf 2 (Tarkett-Sommer® medium pile, sand infill); Turf 3 (FieldTurf® long pile, sand/rubber infill). Turf 1 and 2 used a 10mm rubber underlay. The turfs were installed over the force plate according to manufacturers’ guidelines.  A subject applied a range of loads by jumping from different heights (35cm and 50cm). Two sizes of circular test foot were used: one to represent foot-flat (F1) contact area (20.1mm²) and the other forefoot (F2) contact area (9.5mm²).  These areas were based on a standard UK size 10 shoe. The test rig placed directly on the force plate, using the foot-flat (F1) test foot was used as a control condition.

The placement of the force-link apparatus above the turf could introduce additional errors.  The force-link does not measure forces below the position of the sensing element of the transducer. Acceleration of masses that are below this position will exhibit an inertial force, which will not be measured by the force-link. These masses include that of the displaced turf, as well as the known mass of the test foot (F1 mass= 593g; F2 mass = 287g) and approximately half of the force-link (mass = 46.5g). To calculate the inertial force of the test foot and half of the force-link, a small single-axis Entran EGCS-D1SM accelerometer (Entran, Les Clayes) (mass=12g) was attached to the test foot.  The accelerometer was used on tests with turf 2.

The outputs from the force plate and force-link was sampled at 1080Hz and collected with a Vicon 612 motion analysis system (Oxford Metrics, Oxford). The output from the accelerometer was sampled at 1000Hz and collected on a 4-channel data logger (Biomedical Monitoring, Glasgow). All the data was filtered with a 4th order Butterworth filter with a cut-off filter set at 25Hz, to remove the resonance of the jump rig.

### Results

The errors were calculated as the difference between the peak forces recorded by the force plate to those recorded from the force link, relative to the peak force link output.  Secondly, the mean of the difference between the two outputs over the first 100 samples from initial contact was calculated to provide an indication of the error during the whole impact phase of the jumps performed.

The force-link measured slightly higher peak values than the force-plate, although the relative errors were below 0.6% (Figure 4.12). The difference between the highest observed error (Turf 3, F1) and the control was approximately 0.3%.

The greatest absolute peak error was 29N at a peak force of 5780N on Turf 2 (F2) (Figure 4.13).  The smallest absolute peak error was 0.05N at a peak force of 548N on Turf 1 (F1). The difference in error between the two sizes of contact areas on all three turfs was small (less than 0.15%).  Turf 3 with the F2 test foot produced the highest mean error of 10 (±6.8) N during the impact phase of the jump.

The inertial force calculated from the acceleration of the test foots F1 and F2 on turf 2 were below 1N for all trials and were therefore deemed insignificant.

Figure 4.12 Mean relative error (±s.d.) between peak force plate and force-link outputs

| Test | Applied Load Range (N) | Peak Error Range (N) |
| --- | --- | --- |
| Control (n=8) | 936-3136 | 1.79-11.8 |
| Turf 1 (F1) (n=11) | 548-4188 | 0.05-11.7 |
| Turf 1 (F2) (n=11) | 633-4310 | 2.2-11.0 |
| Turf 2 (F1) (n=11) | 644-5784 | 0.95-29.4 |
| Turf 2 (F2) (n=11) | 678-4076 | 2.02-12.1 |
| Turf 3 (F1) (n=8) | 602-4439 | 1.0-28.1 |
| Turf 3 (F2) (n=8) | 668-3027 | 0.3-16.2 |

Table 4.8 Ranges of dynamic applied loads

Figure 4.13 Mean absolute errors (taken over 100 samples from first impact) (±s.d) between force plate and force-link outputs

### Discussion

This errors calculated in this study could arise from a combination of the inertial effects of the artificial surface and the test foot, and the accuracy of the measuring systems.  Nigg (1990) reported that the inertial effect of artificial turf was approximately 4N, although no data was presented. This study produced higher absolute values at peak forces, although it was considered insignificant when calculated relative to the magnitude of the applied peak force.

It was observed that the higher the applied peak forces, the higher the error between the two outputs. Analysis of the data over the whole impact phase (100 samples/93msec from the initial impact) showed that the mean error magnitude was below 10N. The FieldTurf (turf 3) artificial turf was the ‘softest’ of the three turfs to be tested and exhibited the greatest inertial effects.  A ‘softer’ turf will displace more under load, absorbing some of the energy of the impact.  Therefore, a greater effective mass of turf will be displaced and have a greater inertial force.

The radius and mass of the test foot has been shown to affect the deformation of the surface during an impact (Nigg & Yeadon, 1987). A relatively large area of artificial surface deforming under loading may exhibit a significant inertial force. During this study however, increasing the size of contact area did not greatly affect the error between the two outputs. In addition, it was observed that the inertial effect of the test foots was minimal.

The accuracy of this type of force plate and force-link was specified to be ±0.5% and ±0.2%, respectively. A combined, total error of the whole measuring system is 0.7%. The relative errors measured in this study fall within or very slightly over these limits and can therefore be regarded as insignificant.

The range of forces applied in this study was large, in an attempt to cover the diversity of applied forces observed in sporting activities. However, this study did not include the assessment of shear forces that are applied during sporting activities. These shear forces are usually lower than the vertical forces observed. It could be inferred from the results of this study that the possible inertial effect of the artificial in the horizontal direction would also be minimal.

The errors observed in this study were insignificant. It can be concluded that any inertial effect of artificial turf during the biomechanical assessment of sporting movements on such surfaces is minimal and validates the use of force plates to measure the loading actions at the human/surface interface.