Documentation for Mathematica solver of Haemoperfusion Column Model, perfuse.ma

Contents

1.	Introduction

2.	How to use perfuse.ma to predict the performance of a haemoperfusion column

3.	The mathematical model

4.	The program explained

5.	References

1.  Introduction

The Mathematica notebook perfuse.ma contains a program which can be used to predict the performance of a packed-bed, activated carbon filled, haemoperfusion column.  The program implements a model which is capable of dealing with the removal of solutes which are either in simple aqueous solution, or bound to plasma proteins.

As input, the program requires information about the activated carbon, the column and the flow rate of the perfusing fluid.  Output is given as a plot of fractional clearance (clearance/flow rate) against perfusion time.

The model implemented by the program was developed by Radcliffe (1978).  Brief details of the model are presented later in this document.  Full details can be found in the original reference, or in Radcliffe & Gaylor (1981)

2.  How to use perfuse.ma to predict the performance of a haemoperfusion column

It would be useful for the user to become familiar with the Mathematica software before trying to use perfuse.ma.  Wolfram (1991) provides a good introduction to using Mathematica.  Mathematica sessions are conducted in a “notebook” environment, in which program input and output are all represented simultaneously on the computer screen.

The program, as it stands, is set up to solve for a perfusion with constant inlet concentration of the toxin (i.e. a single pass perfusion).  The carbon in the column is assumed to have no toxin adsorbed to it, prior to the start of perfusion.

For these conditions, the steps to obtain a prediction of the column’s performance are:

1.  Start Mathematica.

2.  Open the Mathematica notebook, perfuse.ma.

3.  Execute the command: Evaluate Notebook.  The program will now begin to run.

4.  Dialogue boxes will appear, asking the user to enter values for the various input parameters required by the program.  The parameters required are listed in Table 1.  Values must be entered for all parameters; any value not provided is likely to be assumed zero.

| Parameter | units |
| --- | --- |
| Simulation name | - |
| Column volume, V | ml |
| Column void fraction, e | - |
| Flow rate, Q | ml/min |
| Carbon granule diameter, dp | mm |
| Carbon bulk density, rb | kg/m3 |
| qm | mmol/g |
| K1 | l/mmol |
| ke | m/s |
| Ds | m2/s |
| Sorbate inlet concentration, co | mg/ml |
| Sorbate molecular weight | g/mol |
| Number of cells to use in model, n | - |
| Protein binding coefficient, K | - |

Table 1.  Input data

5.  After all data has been entered, open the Main Program section.  The numbers being printed on the screen represent the number of completed time steps.

6.  Once solution has proceeded for all time steps, open the Output section to see the graphical results.  A typical output plot from the program is shown in Figure 1.  The plot is in a format which can be exported to other software, such as Microsoft Word or Excel.

Figure 1.  Typical output from perfuse.ma

3.  The mathematical model

perfuse.ma represents a haemoperfusion column according to the lumped parameter model of Radcliffe (1978):

The model considers a column of volume V, with a blood flow rate Q.  The carbon granules have diameter dp, specific surface area ap and pack in the column with a void fraction e.  For the carbon/sorbate combination under consideration, the external mass transfer coefficient is ke, the intraparticle diffusivity is Ds and the Langmuir isotherm parameters are qm and K1.  To account for axial mixing, the column is modelled as n well mixed tanks in series.  The sorbate concentration at inlet is co.  Asterisks denote concentration values at the sorbent/fluid interface.

The normalised model equations for the jth tank are as follows:

dXjdT=Lep(n(Xj−1−Xj)−dYjdT)dYjdT=Ne(Xj−Xj∗)Xj∗Xj−NsyqNeYj∗2−Yj22YjYj∗Xj∗R+(1−R)Xj∗

X=cco		qo=qmK1co1+K1co			R=11+K1co

Y=qqo		Ns=60DsLVdp2Q		L=ρbqoco

T=QLVt		Ne=keap(1−e)VQ

For protein bound sorbates, where the protein binding can be adequately represented by the linear relationship,

Total sorbate concentration = K  x  Free sorbate concentration per unit water volume

the dimensionless parameters in the column model must be modified as follows:

R'=11+1K1R−1, 		 T'=TRKR',

Ns'=NsR'RK,	Ne'=NeK,	L'=LR'KR

The model is solved by a time-marching solution procedure, whereby the solution proceeds in discrete time steps DT.  This is achieved by modifying the column equations as shown:

DXjDT=Le(n(Xj−1−Xj)−DYjDT)                          1DYjDT=Ne(Xj−Xj∗)                                            2Xj∗Xj−NsyqNeYj∗2−Yj22Yj                            3Yj∗Xj∗R+(1−R)Xj∗                                            4

Solution proceeds by the following sequence:

1.  Define initial conditions for Xj and Yj in each cell (j = 1 to n) at dimensionless time T=0

2.  Calculate X1* iteratively from equations {3} and {4}, using the initial values of X1 and Y1.

3.  Calculate DY1/DT from equation {2}, using X1* and the initial values of X1 and Y1.

4.  Calculate the new value of X1 from equation {1}

5.  Calculate the new value of Y1 = (initial value of Y1) + (DY1/DT). DT

6.  Repeat steps 2 to 5 for the remaining cells (j = 2 to n).

7.  Calculate the second time step by repeating steps 2 to 6, using the newly calculated values of Xj and Yj in place of the initial values.

8.  Continue the process for as many time steps as required.

4.  The program explained

This section explains each individual section of the notebook perfuse.ma.  It is not necessary to read this section before using the model, but it may be of interest to anyone who wishes to modify the program.

Initialisation section:  This section ensures that the program runs smoothly.  The code tells Mathematica not to print low-priority warning messages, and also ensures that the variables to be used in the program are not still assigned a value from any previous calculation.

Data input dialogues:  This section contains the coding for the interactive data input dialogue boxes.

Calculate dimensionless parameters:  Calculates the dimensionless parameters required by the model, from the data entered in the previous section.  Parameter names are all chosen to begin with lower case letters, to avoid possible confusion with built-in Mathematica functions, which all begin with upper case letters.

Modify parameters if protein binding:  Allows the model to be used to predict the ability of the haemoperfusion column to remove protein bound solutes. The dimensionless parameters are modified by a factor related to the bound fraction of solute in the solution.

Set initial conditions:  At T=0, both X and Y are zero in all cells of the model.  A step input of X=1 is applied at the inlet to the first cell.

Definition of xstarr(x,y) function:  This section is a subroutine which uses the user defined function, xstarr(x,y) to calculate X* by iteration.

Main loops:  These loops carry out the time marching calculation steps detailed in section 3.

Plot Output:  Provides output as shown in Figure 2.

Manual:  Contains this document.

5.  References

Radcliffe, D. F. (1978):  Mass Transfer in Haemoperfusion Columns and other Sorbent Based Devices for Blood Detoxification.  PhD Thesis, Strathclyde University, 1978

Radcliffe, D. F., Gaylor, J. D. S. (1981a):  Sorption kinetics in haemoperfusion columns. Part 1: estimation of mass transfer parameters.  Med. & Biol. Eng. & Comput., 19, 617-626, 1981.

Radcliffe, D. F., Gaylor, J. D. S. (1981b):  Sorption kinetics in haemoperfusion columns. Part 2: modelling column performance.  Med. & Biol. Eng. & Comput., 19, 627-637, 1981.

Wolfram, S.  (1991):  Mathematica.  Addison Wesley, 1991.